Brüel & Kjær - Noise Dose meter Link Software VP 7790  1.0    © 2000 by Brüel & Kjær

NOISE DOSEMETER SURVEY REPORT FORM

Report format	ISO Noise at Work Regulations
data filename	c:\b&k_app\fhhojt.dta

User entered information:

Company name	
Wearer's name	
Location	
Department	
Job function	
Payroll number	
Social Security number	

User entered notes:

Notes

Setup information:

Setup name	ISO85
Dosemeter model number	Brüel & Kjær-4443
Dosemeter serial number	000000
Frequency weighting for RMS	A
Frequency weighting for Peak	Linear
Time weighting	Fast
Measurement range	70 - 140  dB

User calibration information:

Instrument last checked on	15-06-2005	10:47:52
Microphone serial number	80000758228

Logging times:			Duration
	dd-mm-yyyy	hh:mm:ss	hh:mm:ss,ss
Start of run	05-10-2005	23:57:00
End of run	06-10-2005	00:02:00
Duration of run			00:05:00,00
Duration of pause			00:00:00,00

Results of measurements:

Sound exposure	0,30  Pa2hr
Sound exposure	1067,80  Pa2sec
Daily personal exposure level Lep,d	79,7  dB
Equivalent sound level Leq	99,5  dB  Q = 3
Sound exposure level SEL	124,3  dB  No threshold; No criterion; Q = 3

Threshold level	75  dB	75  dB
Criterion level	85  dB	90  dB
Exchange rate Q	3	3
Actual measured dose	29,32  %	9,2719  %
8 Hour projected dose	2814,74  %	890,10  %
Average sound level Lavg	99,5  dB	99,5  dB
Time weighted average level TWA	79,7  dB	79,7  dB
User specified projected period	08:00  hh:mm	08:00  hh:mm
User specified projected dose	2814,74  %	890,10  %

Exceedance times:
	hh:mm:ss.ss	for % of run
Time under range	00:00:02,00	0,60  %
Above or equal to 70  dB	00:04:58,20	99,40  %
Above or equal to 85  dB	00:04:29,12	89,71  %
Above or equal to 90  dB	00:04:15,19	85,06  %
Above or equal to 100  dB	00:01:49,54	36,51  %
Above or equal to 115  dB	00:00:00,00	0,00  %
Above or equal to 140  dB	00:00:00,00	0,00  %
Time over range	00:00:00,00	0,00  %
Above or equal to user specified level	00:04:46,36	95,45  %
User specified level	80  dB

Maximum exceedance levels:
		at dd-mm-yyyy hh:mm:ss.ss
RMS maximum level (Fast)	108,1  dB	at  06-10-2005 00:01:42
Peak exceedance level	120,1  dB	at  05-10-2005 23:59:47

Statistical exeedance levels:

L10.0%	103,5  dB
L50.0%	97,0  dB
L90.0%	84,5  dB
L95.0%	80,5  dB
L99.0%	72,0  dB

Time history profiles:

Profiles	Off
Profile sample interval	n.a.
Profile function 1	n.a.
Profile function 2	n.a.